Automated Liquid Handler

Technical Design Document

An autonomous robotic liquid handling system for lab environments

System Overview

An automated liquid handler built on a repurposed Sovol SV08 3D printer as the XYZ motion platform. A pipette assembly driven by a stepper linear actuator replaces the print head, enabling precise aspiration and dispensing of liquid samples. An Atlas Scientific EZO pH circuit provides real-time pH monitoring. A Raspberry Pi 5 runs a SILA2 server that coordinates instrument operation and communicates with a desktop client over the local network. An onboard camera provides live monitoring and computer vision feedback.

| Component | Role |
| --- | --- |
| Sovol SV08 | XYZ motion platform, onboard H616 SoC runs Klipper and Moonraker API |
| SV08 MCU | Real-time stepper pulse generation for X, Y, Z, and pipette axes |
| Raspberry Pi 5 | Lab controller, runs SILA2 server, EZO driver, camera CV processing |
| Stepper Linear Actuator | Drives pipette plunger for aspiration and dispensing |
| Pipette Cone Assembly | Liquid displacement mechanism |
| Atlas Scientific EZO pH Circuit | Converts analog probe voltage to digital pH reading over I2C |
| Micro pH Probe | Measures pH of liquid samples |
| Onboard Camera | Live feed and CV-based arm monitoring via Crowsnest |
| Desktop Computer | User interface, connects to SILA2 server on Raspberry Pi |

# Wiring and Pin Out Diagrams

Hardware Overview

XYZ Motion Platform

Overview

The Sovol SV08 CoreXY 3D printer serves as the XYZ motion platform. The CoreXY gantry handles X and Y movement, and the Z axis raises and lowers the pipette assembly over the workspace. Motion is controlled by Klipper firmware. The H616 SoC on the SV08 mainboard runs the Klipper host and exposes the Moonraker REST/WebSocket API for external control. A separate MCU on the same board handles real-time stepper pulse generation.

Components

Sovol SV08

The base motion platform. Modified by replacing the print head with the pipette assembly. Klipper config files are updated to define the liquid handler workspace, safe acceleration limits, and pipette axis.

| Parameter | Value |
| --- | --- |
| Motion system | CoreXY |
| Build volume | 350 x 350 x 345 mm |
| Onboard computer | Allwinner H616 ARM SoC |
| Firmware | Klipper (Sovol fork) |
| API | Moonraker (port 7125) |
| Camera software | Crowsnest (port 8080) |
| Z axes | 4 independent stepper motors |

Raspberry Pi 5

Main lab controller. Connects to the SV08 H616 over the local network via the Moonraker API. Hosts the SILA2 server and drives the Atlas EZO pH circuit over I2C.

| Parameter | Value |
| --- | --- |
| Model | Raspberry Pi 5 |
| RAM | 8 GB |
| OS | Raspberry Pi OS (64-bit) |
| Connection to SV08 | Ethernet, local network |
| Moonraker API port | 7125 |
| Crowsnest stream port | 8080 |
| pH sensor interface | I2C (GPIO pins) |

Liquid Handler

Overview

The print head of the SV08 is replaced with a pipette assembly. A stepper linear actuator drives the pipette plunger to aspirate and dispense liquid. The actuator is controlled as an additional stepper axis through the SV08 firmware. An Atlas Scientific EZO pH circuit connected to the Raspberry Pi provides real-time pH feedback during operation.

Components

Stepper Linear Actuator

Drives the pipette plunger. Controlled via the SV08 stepper driver as an additional axis in Klipper config.

| Parameter | Value |
| --- | --- |
| Manufacturer | Haydon Kerk Pittman |
| Type | Can-stack stepper linear actuator |
| Control path | SV08 stepper driver via Klipper |

Pipette Cone Assembly

| Parameter | Value |
| --- | --- |
| Manufacturer | Pipette Supplies |
| Model | E1 ClipTip single-channel |
| Max volume | 1250 uL |

Atlas Scientific EZO pH Circuit

Converts the analog millivolt signal from the pH probe into a digital pH reading. Communicates with the Raspberry Pi over I2C. An electrically isolated carrier board is used to prevent noise from the stepper motors interfering with readings.

| Parameter | Value |
| --- | --- |
| Manufacturer | Atlas Scientific |
| Model | EZO pH Circuit |
| Interface | I2C |
| I2C address | 0x63 (default) |
| Carrier board | Electrically Isolated EZO Carrier Board Gen 2 |
| Input | Analog mV from pH probe |
| Output | pH value over I2C |

Micro pH Probe

| Parameter | Value |
| --- | --- |
| Model | TBD |
| Type | Micro pH probe |

Software Overview

| Layer | Deployment | Description |
| --- | --- | --- |
| Klipper firmware | SV08 H616 SoC | Motion planning and G-code execution |
| Moonraker | SV08 H616 SoC | REST/WebSocket API exposing Klipper to the network |
| Crowsnest | SV08 H616 SoC | MJPEG camera stream over HTTP |
| Motion driver | Raspberry Pi | Moonraker API client for sending motion and pipette commands |
| pH driver | Raspberry Pi | Atlas EZO I2C driver for reading pH values |
| Camera CV | Raspberry Pi | Consumes Crowsnest stream for arm monitoring |
| SILA2 server | Raspberry Pi | Exposes instrument as a SILA2-compliant lab device |
| Desktop client | Desktop computer | SILA2 client providing user interface and protocol control |

References

[1] Sovol, “SV08 3D Printer,” GitHub. [Online]. Available: https://github.com/Sovol3d/SV08. [Accessed: Apr. 2026].

[2] K. Van and contributors, “Klipper 3D Printer Firmware,” klipper3d.org. [Online]. Available: https://www.klipper3d.org. [Accessed: Apr. 2026].

[3] A. Chacon and contributors, “Moonraker — Klipper API Server,” moonraker.readthedocs.io. [Online]. Available: https://moonraker.readthedocs.io. [Accessed: Apr. 2026].

[4] SiLA Consortium, “SiLA 2 Standard for Lab Automation,” sila-standard.com. [Online]. Available: https://sila-standard.com/standards/. [Accessed: Apr. 2026].

[5] Atlas Scientific, “EZO pH Circuit,” atlas-scientific.com. [Online]. Available: https://atlas-scientific.com/embedded-solutions/ezo-ph-circuit/. [Accessed: Apr. 2026].

[6] M. D. Mikolajczyk et al., “How to convert a 3D printer into a personal automated liquid handler,” SLAS Technology, 2024, doi: 10.1016/j.slast.2024.100173. [Online]. Available: https://www.sciencedirect.com/science/article/pii/S2472630324001213. [Accessed: Apr. 2026].